HANDLING AND OPERATIONAL INSTRUCTIONS

- Store thermistors only in original packaging. Do not open the package prior to storage.

- Storage conditions in original packaging: - storage temperature −40 °C ... +85 °C, relative humidity ≤75%

- Do not store thermistors where they are exposed to heat or direct sunlight. Else the packing material may be deformed, or components may stick together, causing problems during mounting.

- Avoid contamination of thermistor surface during storage, handling and processing.

- Avoid storage of thermistors in harmful environments like corrosive gases such as SOx, Cl etc.

- Use the components within 30 days after opening the factory seal.

- Solder thermistors within the time specified after shipment. For leaded components this is within 24 months.

- NTC thermistor should not be chipped off or any other damage occurred while handling.

- Washing may cause damages to the sensor.

Bending/Twisting

- Bending on wire is permitted at a minimum distance of twice the wire's diameter plus 4 mm from the component. When bending, ensure the wire is mechanically relieved at the component. The bending radius should be at least eight times the wire's diameter.

- Twisting is prohibited as it may cause cracks and or reduce bonding between insulation and coating/potting material.

- Ensure that the materials used are chemically neutral.

- Thermistor Should not overheat during the subsequent assembly process

- Tensile forces on cables or leads should be avoided during the mounting and operation.

- Bending or twisting of cables or leads directly on the thermistor body is not permissible.

- Use thermistors within the specified power range.  Maximum permissible current through the thermistor be 1 mA. For the RTD less than 0.3mA.

Soldering

- Use resin-type flux or non-activated flux.

- Rapid cooling by dipping in solvent is not recommended.

- Complete removal of flux is recommended

Instruction for Disposition

- Disposal shall be carried out in full compliance with the regulatory requirements and environmental guidelines applicable to the respective country.